## 2026-06-02
Laumas TLB4-MODBUSTCP | [6 empty] | Campbell Scientific CR1000Xe
Test Weight |  | Measured |  |  | Error | Error | Test Weight |  | Measured |  |  | Error | Error
lbs | mL equiv | lbs | lbs (zeroed) | mL equiv | mL equiv | (percent) | lbs | mL equiv | lbs | lbs (zeroed) | mL equiv | mL equiv | (percent)
0 | 0 | 0.16763 | 0 | 0 | 0 | #DIV/0! | 0 | 0 | 0.241 | 0 | 0 | 0 | #DIV/0!
0 | 0 | 0.16763 | 0 | 0 | 0 | #DIV/0! | 0 | 0 | 0.241 | 0 | 0 | 0 | #DIV/0!
1.77 | 802.8585 | 1.978 | 1.8104 | 821.1728 | 18.3144 | 2.281% | 1.7 | 771.107 | 1.94 | 1.699 | 770.6534 | -0.45359 | -0.05882%
5.17 | 2345.0726 | 5.3641 | 5.1965 | 2357.0702 | 11.9977 | 0.5116% | 5.21 | 2363.2162 | 5.47 | 5.229 | 2371.8345 | 8.6183 | 0.3647%
11.25 | 5102.9142 | 11.5328 | 11.3651 | 5155.1406 | 52.2265 | 1.023% | 11.6 | 5261.6715 | 11.87 | 11.629 | 5274.8257 | 13.1542 | 0.25%
18.96 | 8600.1113 | 19.3107 | 19.1431 | 8683.1425 | 83.0312 | 0.9655% | 19.14 | 8681.758 | 19.44 | 19.199 | 8708.5199 | 26.7619 | 0.3083%
27.45 | 12451.1106 | 27.8597 | 27.6921 | 12560.9032 | 109.7926 | 0.8818% | 28.16 | 12773.1611 | 28.49 | 28.249 | 12813.5309 | 40.3697 | 0.3161%
34.82 | 15794.0863 | 35.3023 | 35.1347 | 15936.836 | 142.7496 | 0.9038% | 34.68 | 15730.5834 | 35.04 | 34.799 | 15784.5609 | 53.9775 | 0.3431%
39.98 | 18134.623 | 40.5323 | 40.3647 | 18309.1131 | 174.4901 | 0.9622% | 39.63 | 17975.8656 | 39.99 | 39.749 | 18029.8431 | 53.9775 | 0.3003%
44.97 | 20398.0489 | 45.5276 | 45.36 | 20574.9419 | 176.893 | 0.8672% | 44.58 | 20221.1479 | 44.95 | 44.709 | 20279.6613 | 58.5134 | 0.2894%
48.07 | 21804.1852 | 48.7125 | 48.5449 | 22019.5978 | 215.4126 | 0.9879% | 48.07 | 21804.1852 | 48.49 | 48.249 | 21885.3783 | 81.193 | 0.3724%
48.07 | 21804.1852 | 48.7125 | 48.5449 | 22019.5978 | 215.4126 | 0.9879% | 48.07 | 21804.1852 | 48.49 | 48.249 | 21885.3783 | 81.193 | 0.3724%
36.16 | 16401.9001 | 36.6434 | 36.4757 | 16545.1121 | 143.212 | 0.8731% | 35.97 | 16315.7175 | 36.33 | 36.089 | 16369.695 | 53.9775 | 0.3308%
27.5 | 12473.7902 | 27.9267 | 27.7591 | 12591.317 | 117.5268 | 0.9422% | 27.37 | 12414.8232 | 27.72 | 27.479 | 12464.2647 | 49.4416 | 0.3982%
18.62 | 8445.8899 | 19.0425 | 18.8748 | 8561.4873 | 115.5973 | 1.369% | 19.41 | 8804.2279 | 19.71 | 19.469 | 8830.9899 | 26.7619 | 0.304%
10.87 | 4930.5491 | 11.164 | 10.9964 | 4987.8647 | 57.3156 | 1.162% | 10.41 | 4721.8966 | 10.69 | 10.449 | 4739.5867 | 17.6901 | 0.3746%
5.05 | 2290.6415 | 5.297 | 5.1294 | 2326.6564 | 36.0149 | 1.572% | 5.45 | 2472.0784 | 5.71 | 5.469 | 2480.6967 | 8.6183 | 0.3486%
1.77 | 802.8585 | 2.0115 | 1.8439 | 836.3798 | 33.5213 | 4.175% | 1.7 | 771.107 | 1.95 | 1.709 | 775.1894 | 4.0823 | 0.5294%
0 | 0 | 0.20115 | 0.033525 | 15.2069 | 15.2069 | #DIV/0! | 0 | 0 | 0.245 | 0.004 | 1.8144 | 1.8144 | #DIV/0!
0 | 0 | 0.20115 | 0.033525 | 15.2069 | 15.2069 | #DIV/0! | 0 | 0 | 0.244 | 0.003 | 1.3608 | 1.3608 | #DIV/0!
[6 empty] | 1.279% | [6 empty] | 0.3468%
Method:
Test conducted with a 50lb capacity Optima load cell (SN: 2210004) connected to the Laumas load cell transmitter. The mV output was read from the transmitter over Modbus and the excitation voltage was considered to be stable and constant at 4.973V (measured at the start of the test with a Fluke 115 multimeter). Conversion from mV to lbs was done using the following equation:
Where:
= Full Scale Load of the load cell
= is the mV per Volt Excitation output of load cell
= is the excitation voltage
= is the measured millivolt output
Areas for Improvement:
The results from this calibration could be used to create an equation that fits better than the one given above, which relies on the load cell manufacturer's specified mV/V sensitivity.
Method:
Test conducted with a 50lb capacity Optima load cell (SN: 2210004) connected to the Laumas load cell transmitter. The mV output was read from the transmitter over Modbus and the excitation voltage was considered to be stable and constant at 4.973V (measured at the start of the test with a Fluke 115 multimeter). Conversion from mV to lbs was done using the following equation:
𝐿𝑜𝑎𝑑=𝐿_𝐹𝑆/(𝑚𝑉𝑉∗𝑉_𝑥 ) ∗𝑚𝑉_𝑚𝑒𝑎𝑠𝑢𝑟𝑒𝑑
Where:
𝐿_𝐹𝑆 = Full Scale Load of the load cell
𝑚𝑉𝑉 = is the mV per Volt Excitation output of load cell
𝑉_𝑥 = is the excitation voltage
𝑚𝑉_𝑚𝑒𝑎𝑠𝑢𝑟𝑒𝑑 = is the measured millivolt output
Areas for Improvement:
The results from this calibration could be used to create an equation that fits better than the one given above, which relies on the load cell manufacturer's specified mV/V sensitivity.
Method:
Test conducted with a 50lb capacity Optima load cell (SN: 2210004) connected to a Campbell Scientific CR1000Xe datalogger. Excitation for the load cell was provided by the Vx2 port on the datalogger and the BrFull instruction in CRBasic was used to measure the loadcell on differential input channel 3. The output of this instruction is the mV output of the load cell divided by the excitation voltage. Conversion to lbs was done using the following equation:
Where:
= Full Scale Load of the load cell
= is the mV per Volt Excitation output of load cell
= is the measured mV output of the load cell divided by the excitation voltage.
Areas for Improvement:
The results from this calibration could be used to create an equation that fits better than the one given above, which relies on the load cell manufacturer's specified mV/V sensitivity.
Method:
Test conducted with a 50lb capacity Optima load cell (SN: 2210004) connected to a Campbell Scientific CR1000Xe datalogger. Excitation for the load cell was provided by the Vx2 port on the datalogger and the BrFull instruction in CRBasic was used to measure the loadcell on differential input channel 3. The output of this instruction is the mV output of the load cell divided by the excitation voltage. Conversion to lbs was done using the following equation:
𝐿𝑜𝑎𝑑=𝐿_𝐹𝑆/𝑚𝑉𝑉 ∗𝑋
Where:
𝐿_𝐹𝑆 = Full Scale Load of the load cell
𝑚𝑉𝑉 = is the mV per Volt Excitation output of load cell
𝑋 = is the measured mV output of the load cell divided by the excitation voltage.
Areas for Improvement:
The results from this calibration could be used to create an equation that fits better than the one given above, which relies on the load cell manufacturer's specified mV/V sensitivity.

## 2026-06-08
DATA FROM JUNE 1, 2026 (Used for Equation) | [5 empty] | TEST DATA FROM JUNE 8, 2026
Test Weight |  | Measured |  | Equation (using LINEST) |  | Test Weight |  | Measured |  |  | Error
lbs | mL equiv | mV | mV zeroed | Slope | Offset | lbs | mL equiv | mV | mV zeroed | mL equiv | mL equiv | %
0 | 0 | 0.05 | 0 | 1506.3586 | 0 | 0 | 0 | 0.06 | 0 | 0 | 0 | 0%
0 | 0 | 0.05 | 0 | Note: I forced the intercept to be 0. I think this is the right idea, but maybe I'm just skewing the data? |  | 0 | 0 | 0.06 | 0 | 0 | 0 | 0%
1.77 | 802.8585 | 0.59 | 0.54 |  |  | 1.8 | 816.4663 | 0.6 | 0.54 | 813.4336 | 3.0326 | 0.3714%
5.17 | 2345.0726 | 1.6 | 1.55 |  |  | 5.76 | 2612.6921 | 1.8 | 1.74 | 2621.064 | 8.3719 | 0.3204%
11.25 | 5102.9142 | 3.44 | 3.39 |  |  | 10.78 | 4889.7257 | 3.3 | 3.24 | 4880.6019 | 9.1239 | 0.1866%
18.96 | 8600.1113 | 5.76 | 5.71 |  |  | 16.84 | 7638.4955 | 5.13 | 5.07 | 7637.2381 | 1.2574 | 0.01646%
27.45 | 12451.1106 | 8.31 | 8.26 |  |  | 23.66 | 10731.9955 | 7.18 | 7.12 | 10725.2733 | 6.7222 | 0.06264%
34.82 | 15794.0863 | 10.53 | 10.48 |  |  | 36.89 | 16733.0225 | 11.16 | 11.1 | 16720.5805 | 12.442 | 0.07436%
39.98 | 18134.623 | 12.09 | 12.04 |  |  | 43.85 | 19890.0254 | 13.25 | 13.19 | 19868.87 | 21.1555 | 0.1064%
44.97 | 20398.0489 | 13.58 | 13.53 |  |  | 48.11 | 21822.3289 | 14.54 | 14.48 | 21812.0726 | 10.2563 | 0.047%
48.07 | 21804.1852 | 14.53 | 14.48 |  |  | 40.81 | 18511.1046 | 12.33 | 12.27 | 18483.0201 | 28.0846 | 0.1517%
48.07 | 21804.1852 | 14.53 | 14.48 |  |  | 31.86 | 14451.4529 | 9.65 | 9.59 | 14445.979 | 5.4739 | 0.03788%
36.16 | 16401.9001 | 10.93 | 10.88 |  |  | 20.08 | 9108.1348 | 6.1 | 6.04 | 9098.406 | 9.7288 | 0.1068%
27.5 | 12473.7902 | 8.33 | 8.28 |  |  | 13.26 | 6014.6348 | 4.05 | 3.99 | 6010.3708 | 4.264 | 0.07089%
18.62 | 8445.8899 | 5.68 | 5.63 |  |  | 8.3 | 3764.8167 | 2.57 | 2.51 | 3780.9601 | 16.1434 | 0.4288%
10.87 | 4930.5491 | 3.33 | 3.28 |  |  | 5.06 | 2295.1774 | 1.58 | 1.52 | 2289.6651 | 5.5123 | 0.2402%
5.05 | 2290.6415 | 1.58 | 1.53 |  |  | 1.8 | 816.4663 | 0.6 | 0.54 | 813.4336 | 3.0326 | 0.3714%
1.77 | 802.8585 | 0.6 | 0.55 |  |  | 0 | 0 | 0.06 | 0 | 0 | 0 | 0%
0 | 0 | 0.06 | 0.01 |  |  | 0 | 0 | 0.06 | 0 | 0 | 0 | 0%
0 | 0 | 0.06 | 0.01 | [6 empty] | Max | 28.0846 | 0.4288%
[10 empty] | Min | 0 | 0%
[10 empty] | Avg | 8.2231 | 0.1439%